# Test of Eye-One display calibrator

A test and evaluation of a display calibrator, Eye-One was performed. A LCD some years old with no preset colour temperature setting but with user settings of red, green and blue and a reset to factory settings was used for the test. The screen was measured with a newly calibrated spectroradiometer from PhotoResearch, PR650. The Eye-One calibrator works fine as long as the start values of the display are not too long from the target values of the calibration. The most important parameters calibrated by the Eye-One are the colour temperature and the gamma. The luminance setting of the Eye-One is set manually with the display brightness control and depends on the operator. It is important to notice that the basic RGB values are not possible to alter because they are determined completely by the colour filters and the background illumination. If the colour gamut is small, it is presently impossible to increase. However, if the background illumination can be modulated, for instance by the use of multicolour LEDs, this problem can be solved. Other parameters not calibrated by the Eye-One is luminance uniformity, and colour uniformity. A good advice is to set the user RGB at the default setting or at rather high values as start values in order to utilize an optimized luminance range. The Eye-One is excellent for displays with no preset colour temperature settings. The Eye-One does not support colour temperatures above 7500K. However, displays complying with the TCO´03 should be accurate enough by them selves and do not need recalibration by the Eye-One calibrator.

Comments to the test and results.

The luminance  reading should not be taken too serious because it is calibrated by hand with the display controls. The colour temperatures are close to the correct value within ± 0.005 units of u´ and v´. The greyscale colour shifts are even better and are within ±0.003 units of u´ and v´. This would be hard to notice for most people.

| Testing of Eye-One display calibrator 20050906 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | Luminance |  |  | CCT |  |  |  |
|  |  | cd/m2 | u´ | v´ | Kelvins |  |  |  |
|  | Start | 134 | 0.198 | 0.481 | 5840 |  |  |  |
|  | Gamma | 2.2 |  |  |  |  |  |  |
|  |  |  |  |  |  | Luminance |  |  |
|  | Tested display | x | y | u´ | v´ | cd/m2 |  |  |
|  | Red | 0.6362 | 0.3424 | 0.4360 | 0.5280 | 30.2 |  |  |
|  | Green | 0.3000 | 0.6065 | 0.1240 | 0.5640 | 69.6 |  |  |
|  | Blue | 0.1404 | 0.0984 | 0.1440 | 0.2270 | 15.4 |  |  |
|  | White (6500K) | 0.312 | 0.319 | 0.201 | 0.463 | 117.0 |  |  |
|  |  | x | y | u´ | v´ |  |  |  |
|  | sRGB red | 0.6400 | 0.3300 | 0.4507 | 0.5229 |  |  |  |
|  | sRGB green | 0.3000 | 0.6000 | 0.1250 | 0.5625 |  |  |  |
|  | sRGB blue | 0.1500 | 0.0600 | 0.1754 | 0.1579 |  |  |  |
|  | sRGB white | 0.3127 | 0.3291 | 0.1978 | 0.4684 |  |  |  |

|  | CCT | Luminance | CIE 1931 | CIE 1976 | CCT | 106/CCT |
| --- | --- | --- | --- | --- | --- | --- |
|  | 6500 K | cd/m2 | x | y | u´ | v´ | Kelvins |  |
|  | Correct | 100 | 0.3126 | 0.3284 | 0.198 | 0.468 | 6500 | 153.85 |
|  | Eye-One | 106.1 |  |  |  |  | 6500 | 153.85 |
|  | 255 | 113 | 0.314 | 0.323 | 0.201 | 0.465 | 6493 | 154.01 |
|  | 225 | 81.5 | 0.315 | 0.324 | 0.201 | 0.466 | 6464 | 154.70 |
|  | 195 | 63.3 | 0.316 | 0.327 | 0.201 | 0.468 | 6299 | 158.76 |
|  | 165 | 43.8 | 0.318 | 0.325 | 0.203 | 0.467 | 6236 | 160.36 |
|  | 135 | 28.8 | 0.320 | 0.323 | 0.205 | 0.466 | 6174 | 161.97 |
|  | 105 | 16.7 | 0.320 | 0.324 | 0.205 | 0.467 | 6128 | 163.19 |
|  | Error CCT |  | -0.001 | 0.006 | 0.003 | 0.003 |  |  |
|  | Error grey |  | 0.007 | 0.005 | 0.004 | 0.002 |  |  |

|  | CCT | Luminance | CIE 1931 | CIE 1976 | CCT | 106/CCT |
| --- | --- | --- | --- | --- | --- | --- |
|  | 7500 K | cd/m2 | x | y | u´ | v´ | Kelvins |  |
|  | Correct | 140 | 0.300 | 0.315 | 0.194 | 0.459 | 7500 | 133.33 |
|  | Eye-One | 125.9 |  |  |  |  | 7500 | 133.33 |
|  | 255 | 141 | 0.296 | 0.311 | 0.193 | 0.456 | 7712 | 129.67 |
|  | 225 | 104 | 0.299 | 0.314 | 0.194 | 0.458 | 7549 | 132.47 |
|  | 195 | 77 | 0.297 | 0.313 | 0.193 | 0.457 | 7654 | 130.65 |
|  | 165 | 54.2 | 0.297 | 0.313 | 0.193 | 0.457 | 7640 | 130.89 |
|  | 135 | 35 | 0.299 | 0.314 | 0.194 | 0.458 | 7505 | 133.24 |
|  | 105 | 20.5 | 0.299 | 0.314 | 0.194 | 0.458 | 7531 | 132.78 |
|  | Error CCT |  | 0.004 | 0.004 | 0.001 | 0.003 |  |  |
|  | Error grey |  | 0.003 | 0.003 | 0.01 | 0.002 |  |  |
|  | CCT | Luminance | CIE 1931 | CIE 1976 | CCT | 106/CCT |
|  | 5000 K | cd/m2 | x | y | u´ | v´ | Kelvins |  |
|  | Correct | 100 | 0.345 | 0.358 | 0.209 | 0.488 | 5000 | 200.00 |
|  | Eye-One | 104.7 |  |  |  |  | 5000 | 200.00 |
|  | 255 | 110 | 0.342 | 0.355 | 0.208 | 0.486 | 5118 | 195.39 |
|  | 225 | 84.4 | 0.341 | 0.362 | 0.205 | 0.489 | 5165 | 193.61 |
|  | 195 | 60.6 | 0.341 | 0.362 | 0.205 | 0.489 | 5203 | 192.20 |
|  | 165 | 43.6 | 0.342 | 0.364 | 0.205 | 0.490 | 5154 | 194.02 |
|  | 135 | 28.8 | 0.338 | 0.365 | 0.202 | 0.490 | 5300 | 188.68 |
|  | 105 | 16.3 | 0.339 | 0.363 | 0.203 | 0.489 | 5270 | 189.75 |
|  | Error CCT |  | 0.003 | 0.003 | 0.001 | 0.003 |  |  |
|  | Error grey |  | 0.004 | 0.010 | 0.01 | 0.002 |  |  |

|  | CCT | Luminance | CIE 1931 | CIE 1976 | CCT | 106/CCT |
| --- | --- | --- | --- | --- | --- | --- |
|  | Native | cd/m2 | x | y | u´ | v´ | Kelvins |  |
|  | Correct |  |  |  |  |  | Native |  |
|  | Eye-One | 182.8 |  |  |  |  | 5800 | 172.41 |
|  | 255 | 186.0 | 0.325 | 0.352 | 0.198 | 0.482 | 5792 | 172.65 |
|  | 225 | 139.0 | 0.327 | 0.352 | 0.199 | 0.482 | 5737 | 174.31 |
|  | 195 | 104.0 | 0.329 | 0.353 | 0.200 | 0.483 | 5630 | 177.62 |
|  | 165 | 73.5 | 0.330 | 0.355 | 0.200 | 0.484 | 5585 | 179.05 |
|  | 135 | 48.5 | 0.330 | 0.353 | 0.201 | 0.483 | 5600 | 178.57 |
|  | 105 | 27.0 | 0.332 | 0.354 | 0.202 | 0.484 | 5499 | 181.85 |
|  | Error grey |  | 0.007 | 0.003 | 0.004 | 0.004 |  |  |